Republic of Latvia

Cabinet

Regulation No. 97

Adopted 5 March 2002

Regulations Regarding Protection Against Ionising Radiation in Medical Exposure

Issued pursuant to Section 3, Paragraph three

of the Law on Radiation Safety and Nuclear Safety

1. General Provisions

1. These Regulations prescribe the implementation of basic principles for radiation safety and nuclear safety in the protection of human beings against ionising radiation in medical exposure.

2. Medical exposure is a process during which ionising radiation for medical purposes affects human tissues, organs and the organism as a whole.

3. These Regulations apply to the following types of medical exposure:

3.1. exposure of persons that is related to the diagnosis or treatment of a disease, or health surveillance;

3.2. exposure of employees in mandatory health surveillance of employees;

3.3. exposure of such persons as voluntarily participate in medical or biomedical research programmes;

3.4. exposure of persons in court medical expert-examinations and health insurance procedures, if the relevant examinations have medical indications; and

3.5. exposure of such persons who come into contact with patients undergoing medical exposure, which persons knowingly and willingly provide support and comfort to such patients if it is not a part of the professional duties of such persons (hereinafter - voluntary assistant).

4. The following radiological procedures, during which human beings are subjected to medical exposure, shall be utilised in radiology:

4.1. radio diagnostic procedures - medical examination utilising a source of radio-nuclides or other ionising radiation only for diagnostic purposes (for example, diagnostic nuclear medicine procedures, roentgenology, computed tomography, invasive radiology, and dental radiology); and

4.2. radiotherapeutic procedures – medical treatment utilising a source of radio-nuclides or other ionising radiation (for example, therapeutic procedures of nuclear medicine, utilisation of accelerators and other equipment containing a source of ionising radiation in radiotherapy and invasive radiology).

5. Medical treatment practitioners who have the right to perform medical radiological procedures (hereinafter - practitioners) shall be liable for the selection of a medical radiological procedure and for each medical exposure. Practitioners have the duty to:

5.1. evaluate the justification for medical exposure;

5.2. ensure the optimisation of medical exposure:

5.2.1. select installations or equipment that is used in radiology and conforms to the radiological procedures, which installations or equipment contain a source of ionising radiation or generate ionising radiation (hereinafter - radiological installations);

5.2.2. obtain the radio diagnostic information or provide a radiotherapeutic result;

5.2.3. evaluate the technical aspects of medical exposure; and

5.2.4. taking into account economic and social factors, to evaluate medical radiological procedures, as well as the dose of ionising radiation or the amount of total radioactivity that a person receives during medical exposure (hereinafter - patient dose);

5.3. evaluate the radio diagnostic information or radiotherapeutic result obtained during medical exposure;

5.4. co-operate with other specialists related to medical exposure with respect to the technical aspects of medical exposure (for example, the preparation and performance of medical radiological procedures, the handling and use of radiological installations related to such procedures, as well as the assessment of the technical parameters of radiological installations and the physical parameters of the workplace, determination of the patient dose using direct or indirect methods of measuring (hereinafter - patient dosimetry), the calibration and maintenance of radiological installations, the development of films, the preparation and use of radioactive materials);

5.5. obtain information regarding previous medical radiological procedures undergone by the person;

5.6. deliver the information to other medical treatment practitioners regarding the medical exposure of a person;

5.7. provide the patient and the voluntary assistant with information regarding the potentially harmful effects of ionising radiation; and

5.8. if a patient has been subjected to a radiological procedure of nuclear medicine, to provide visitors to the patient with information regarding the potentially harmful effects of ionising radiation.

6. The non-profit-making organisation State stock company "Veselības statistikas un medicīnas tehnoloģiju aģentūra" [Health Statistics and Medical Technology Agency] (hereinafter - agency) shall:

6.1. prepare recommendations regarding the required quantity of radiological installations, compile information regarding the utilisation of medical radiological procedures and alternative methods in the State, and jointly with Radiation Safety Centre (hereinafter - centre) promote compliance with the requirements for ensuring the requirements of radiation safety quality and the requirements for quality control; and

6.2. compile information regarding the medical radiological procedures performed in the State and provide such information to the centre so that the centre can assess the average patient dose in various medical radiological procedures and in groups of patients, and can include the information referred to in a database regarding the exposure of the population.

2. Justification for Medical Exposure

7. A medical treatment practitioner who has the right to send a person for medical exposure (hereinafter - prescriber) in conformity with his or her professional activities, shall complete the referral form by justifying the necessity for medical exposure.

8. A practitioner shall evaluate the justification for medical exposure provided by the prescriber and, if such justification is sufficient, shall permit the medical exposure to be performed.

9. Medical exposure without relevant justification is prohibited.

10. In the justification and evaluation of the necessity of medical exposure:

10.1. the possibilities for curing a patient, reducing the expected pain, increasing medical treatment possibilities or lengthening of the life span and public benefit shall be compared with the possible harm created by the effects of ionising radiation, also taking into account the efficiency, benefit and risk presented by alternative methods, which are not related to or are less related to such medical exposure (for example, ultrasonic examination and chemotherapy);

10.2. the goals of medical exposure and the physical characteristics of a person shall be taken into account;

10.3. the reference levels of the dose of ionising radiation in radio diagnostic procedures, which levels are specified in Annexes 1, 2, 3, 4, 5, 6 and 7 of these Regulations, or - with respect to radioactive pharmaceutical preparations - reference levels of total radioactivity for the examination of standard-sized patients or for standard examinations, utilising radiological installations in accordance with the relevant description of the method utilised in medical exposure (hereinafter - diagnostic reference levels) shall be taken into account. Radioactive pharmaceutical preparations shall be radio-nuclides or isotope-marked chemical compounds, which are intended for administering into a human body or for in vitro activities for diagnostic or therapeutic purposes in nuclear medicine; and

10.4. the previously obtained radio diagnostic information or the data of health history that is related to the planned medical exposure shall be taken into account.

11. Practitioners, in justifying the medical exposure of a voluntary assistant, shall take into account the direct benefit to the health of the patient and the potential harm caused by such exposure to the voluntary assistant or the patient.

12. Practitioners shall evaluate the justification for medical exposure in detail:

12.1. if a new medical radiological procedure is utilised before the implementation of such procedure in the medical treatment institution;

12.2. if evidence of insufficient efficiency of the relevant medical radiological procedure or possible adverse side effects have been acquired; or

12.3. if there is no direct benefit to the health of the person from the medical exposure (also if the person himself or herself requests such exposure). In such cases, medical exposure may be carried out taking into account the restrictions determined by the practitioner, except the case when such exposure causes threat to the health or life of the person.

13. If the utilisation of a medical radiological procedure is not justified in general, in individual cases doctors' council may take a decision regarding an individual medical exposure.

14. The agency in co-operation with the Radiation Safety Council shall evaluate the feasibility and possibility of the implementation of a new medical radiological procedure in the State, by analysing the impact of such procedure on radiation safety and evaluating whether the positive result shall exceed the total negative impact.

15. If the clinical manifestations of a disease, for the medical treatment of which medical radiological procedure is necessary, have not been determined for a person, such procedure in the court medical expert-examination or in health insurance shall be justified only if after such procedure it is possible to obtain the information regarding the state of health of the person who is examined, which information may be utilised in the medical treatment of the person.

16. Medical exposure of large groups of the population (more than 1000 persons) for radio diagnostic purposes, including x-ray radioscopic and mammographic examinations (hereinafter - screening examinations), is only permitted if the anticipated benefit for the relevant group of population or for the whole population compensates for the economic and social expenditures that may be caused by the harmful effects of the ionising radiation, and if the opportunities of such exposure in discovering of diseases, medical treatment possibilities of the discovered diseases and public benefit from the restriction of the spread of such diseases have been taken into account. The necessity of such medical exposure shall be evaluated by the agency.

17. Persons may not undergo exposure in a medical or biomedical research (hereinafter - research), except in a case, where consent for the commencement of a research programme has been obtained from the Latvian Academy of Sciences Ethics Commission (hereinafter - Ethics Commission) and the person to be exposed is informed regarding the expected risk and has given written consent to participate in such research. The Ethics Commission shall supervise that the number of persons involved in research and subjected to medical exposure shall be as small as possible, but sufficient in order to obtain precise and reliable information.

18. Before a patient (a sick person) undergoes exposure in a research, it shall be necessary to evaluate the possible diagnostic and therapeutic benefit that the patient (the sick person) could gain by participating in such research, and to obtain written consent from his or her attending practitioner.

19. The Ethics Commission shall examine prior to the taking a decision regarding the exposure of a person in a research, whether the person intended to be involved in the research and subjected to medical exposure:

19.1. is not dependent on the person performing the research; or

19.2. has not given consent to participate in the research for payment, except the direct compensation for the participation in the research, or has been otherwise influenced.

20. The person performing a research shall be liable for:

20.1. the person intended to be involved in the research and subjected to medical exposure at least one day prior to the exposure is informed regarding the potential risk in writing in a way that he or she can comprehend;

20.2. the person intended to be involved in the research and subjected to medical exposure receives the information regarding the potential effects of the research on the health, as well as regarding the side effects that may occur during the research or afterwards;

20.3. the confidentiality of the information regarding the persons involved in the research is ensured; and

20.4. the persons involved in the research are provided with the opportunity to decline participation in the research at any time, without providing justification for such refusal.

3. Duties of Operators in Radiology

21. Operators shall ensure the development of a programme for radiation safety and nuclear safety quality assurance (hereinafter - quality assurance programme), by inviting a practitioner, radiation safety expert (or radiation safety and nuclear safety expert) and medical physics expert having an education in medical physics and at least three years experience in ionising radiation physics or in the technology of ionising radiation, which technology is utilised in medical exposure, and the competence of which medical physics expert has been recognised by a decision of the radiation safety and nuclear safety expert certification commission.

22. In developing a quality assurance programme, an operator shall include the following activities in such programme:

22.1. to determine the source of ionising radiation, the measurements of technical parameters of the equipment for obtaining the image, of the facilities for the protection against ionising radiation and of the technical equipment related to the generating of ionising radiation, which measurements are necessary upon putting the relevant equipment into service and in the performance of regular examinations;

22.2. to ensure that medical exposure is not performed by such radiological installations the technical parameters of which do not conform to the minimum criteria for technical parameters of radiological installations (Annex 8);

22.3. to determine the mutual conformity of radio diagnostic and radiotherapeutic clinical tasks with the technical aspects of medical exposure;

22.4. to prepare the description of the method to be utilised in medical exposure for all radiological installations and all medical radiological standard procedures;

22.5. to regularly control patient dosimetry and examine the conformity of the operation of measuring instruments with the calibration data;

22.6. to examine measuring instruments daily by utilising standards of work or other measuring equipment and data regarding the operating conditions of such measuring instruments, as well as to examine the medical exposure control equipment;

22.7. to perform regular quality control examinations in radiology by involving the experts of related institutions;

22.8. to perform special quality control measures with respect to the medical exposure of children and the evaluation of patient dose or medical radiological procedures performed on children; and

22.9. to take into account indicators for evaluation of the quality of medical exposure and requirements for the quality assurance programme prescribed by other regulatory enactments.

23. Operators shall be liable for:

23.1. a systematic audit of the quality of radiology, which audit includes the examination of medical radiological procedures by evaluating medical radiological procedures and comparing the results obtained with the diagnostic reference levels, if necessary, by transforming medical radiological procedures, developing the description of the new method to be utilised in medical exposure or by utilising other methods, in order to improve the quality of patient care;

23.2. the conformity of the minimum workload of medical physics experts and medical physicists involved in the work of the medical treatment institution with the requirements specified in Annex 9 of these Regulations;

23.3. the involvement of medical physics experts and medical physicists in radiotherapeutic procedures, and in diagnostic and therapeutic procedures of nuclear medicine. Medical physics experts, if necessary, shall be involved also in other medical radiological procedures;

23.4. placement of patients and employees who are related to medical exposure in premises that conform to radiation safety requirements;

23.5. the recording, collection, labelling, storing for degradation storage or removal into the environment of such radioactive waste (for example, excretion of patients), radioactively contaminated materials and reusable objects as are created from patients or in relation to patients, on which radiological procedures of nuclear medicine are or have been performed;

23.6. regular dosimetry measurements in premises where radiological procedures of nuclear medicine are performed or in which patients, on whom radiological procedures of nuclear medicine are or have been performed, are situated;

23.7. access by voluntary assistants to radiation safety instructions; and

23.8. medical exposure in conformity with the description of the method utilised in medical exposure.

4. Medical Radiological Procedures

4.1. General Requirements

24. Depending on the purpose of exposure and the organ or tissues to be exposed, organs that are sensitive to the ionising radiation (organs of lesser pelvis, mammary glands and thyroid glands) shall be protected during medical radiological procedures.

25. The dose of ionising radiation received by a visitor from the patient who is or has been subjected to the radiological procedure of nuclear medicine, may not exceed 5 mSv per year for the whole body. If the visitor has not reached 18 years of age, the received dose of ionising radiation may not exceed 1 mSv per year for the whole body.

26. Practitioners shall ensure that the dose of ionising radiation received by voluntary assistants does not exceed 5 mSv per year.

27. If the dose of ionising radiation assessed for the voluntary assistant may exceed 5 mSv per year, such assistant shall be equated to an employee who performs activities with the sources of ionising radiation, and the operator shall provide individual dosimetry for such assistant.

28. If a patient who has been subjected to the radiological procedure of nuclear medicine leaves the medical treatment institution, the practitioner shall issue to the patient or to the member of his or her family, trustee or guardian written instructions regarding the measures to be taken for the protection against ionising radiation and information regarding the potential dose of ionising radiation that other persons may receive from the relevant patient.

4.2. Medical Exposure for Radio diagnostic Purposes

29. In radio diagnostic procedures the practitioner shall be liable for:

29.1. utilisation of appropriate radio diagnostic installations;

29.2. exposure of patients to a possibly smaller dose of medical exposure, taking into account economic and social factors and choosing such parameters of medical exposure as ensure the obtaining of a qualitative image;

29.3. not exceeding the diagnostic reference levels and taking into account the technical parameters of the relevant radiological installation; and

29.4. the possibility to avoid additional medical exposure by utilising as much as possible the radio diagnostic information obtained in the previous medical exposure.

30. In order to reduce the patient dose and concurrently provide the necessary radio diagnostic information, the following parameters of exposure shall be utilised in radio diagnostic procedures:

30.1. size of the field to be examined;

30.2. the number and size of digital images, and the number of pictures during the examination or the number sections in computed tomography;

30.3. type of image detector, including the intensifying screen of high sensitivity or low sensitivity;

30.4. detection grid of dispersed ionising radiation;

30.5. collimation of the primary beam;

30.6. the technique for the improvement of images in dynamic imaging, including the number of images per second;

30.7. the temperature of the developer, algorithm for the reconstruction of images and other factors related to the treatment of images; and

30.8. the high voltage supplied to the x-ray tube, electric power, duration of examination and other technical parameters of radiological installations.

31. Mobile radiological installations shall be utilised in radiological procedures only if due to medical contra-indications it is not possible to transfer the patient to the radio diagnostic room.

32. Practitioners in diagnostic procedures of nuclear medicine shall:

32.1. select a radionuclide with a short half-life;

32.2. select such total radioactivity of the dosage to be administered and of the algorithm for the obtaining of images, so that the combination thereof would ensure minimum patient dose and a qualitative image;

32.3. take into account the diagnostic reference levels of total radioactivity;

32.4. comply with special requirements for patients with the dysfunction of organs or systems of organs; and

32.5. utilise methods that accelerate the removal of pharmaceutical radioactive preparations.

33. If a diagnostic procedure of nuclear medicine is necessary for a child, the practitioner shall make adjustments to the total radioactivity of the dosage to be administered to the patient by taking into account the physical parameters of the child, and shall choose an appropriate amount between the minimum dose of a radioactive pharmaceutical preparation for children (Annex 7) and the maximum dose for adults (Annex 6).

34. Practitioners shall be liable for the radioactive pharmaceutical preparation to be used in a diagnostic procedure of nuclear medicine is prepared, the total radioactivity of the dosage to be administered to the patient is measured and the amount of such dose is recorded directly before the medical radiological procedure is performed.

4.3. Medical Exposure for Radiotherapeutic Purposes

35. Practitioners shall plan the medical exposure of the relevant part of the body in radiotherapeutic procedures individually, so that the dose of ionising radiation on the surrounding tissues, which are not the objects of medical exposure, would be as low as reasonably achievable in conformity with the intended radiotherapeutic objective.

36. Practitioners in therapeutic procedures of nuclear medicine shall:

36.1. select a radio nuclide that conforms to the radiotherapeutic procedure; and

36.2. comply with the special requirements for patients with dysfunction of organs or systems of organs.

37. Practitioners shall be liable for:

37.1. utilisation of appropriate radiological equipment in radiotherapeutic procedures;

37.2. receiving the necessary dose of ionising radiation by the planned target volume, protecting other organs and tissues as much as possible; and

37.3. preparation of the radioactive pharmaceutical preparation to be used in a therapeutic procedure of nuclear medicine, measurement of the total radioactivity of the dosage to be administered to the patient and recording of the amount of such dose directly before the medical radiological procedure is performed.

38. Practitioners shall place the patients who have undergone a therapeutic procedure of nuclear medicine in a special ward, in order to protect employees related to medical exposure, other patients and visitors against the ionising radiation.

39. In order to limit persons receiving radiation from patients who have undergone therapeutic procedure of nuclear medicine, the operator shall ensure that the patient is not discharged and released from the medical treatment institution until the total radioactivity present in his or her body is less than 400 MBq 131I, but if the patient lives together (in one room) with minor children - 200 MBq 131I. If other radionuclides are used, the total level of radioactivity shall be determined by the practitioner, taking into account the protection against the ionising radiation of the persons with whom the patient lives or works together.

5. Training

40. In order to increase the level of radiation safety in the State, the centre in co-operation with the Association of Roentgenologists and Radiologists of Latvia shall develop recommendations regarding curricula related to radiation safety in medical exposure that is to be included in training programmes.

41. Persons acquiring relevant training programmes may perform medical radiological procedures (or parts thereof) only under the supervision or management of certified medical treatment practitioners or medical physicists in conformity with the specifications of the work to be performed.

42. Operators shall provide the employees related to medical exposure with:

42.1. training after the qualification is acquired and with the testing of knowledge regarding medical radiological procedures and protection against ionising radiation and regarding other aspects of radiation safety at least once per year, but the work supervisor with training and testing of knowledge - at least once every three years; and

42.2. extracurricular training and testing of knowledge if the regulatory enactments regulating the issues of radiation and nuclear safety are substantially amended, the duties of work or office are changed, or a new radiological installation or a new medical radiological procedure is implemented.

6. Radiological Installations

43. Operators shall ensure that the radiological installations used in radiodiagnostic procedures and the equipment thereof conform to the following requirements:

43.1. installations and the equipment thereof shall be designed and manufactured so as to maintain medical exposure in conformity with diagnostic reference levels, that would allow to obtain sufficient radiodiagnostic information;

43.2. high-voltage of the generating tube, electric power and exposition time or the electric power multiplied by the exposure time is clearly and accurately indicated on the control panel;

43.3. easy and uninterrupted access to the information regarding filtration, size of the focus field, distance from the source of ionising radiation to the image detector, indication of the size of the field, as well as information regarding other parameters that characterise the operation of radiological installations;

43.4. installation shall be equipped with a device that after a specified period of time or after reaching the given dose of ionising radiation automatically discontinues generating the ionising radiation; and

43.5. installation and the equipment thereof shall be appropriate for the performance of medical exposure in health examinations, and for the performance of such medical exposure in which large patient doses are received, including invasive radiology and computed tomography.

44. Fluoroscopic examinations are prohibited:

44.1. without an electronic optical image intensification device or equivalent technology; or

44.2. in screening examination programmes - without a device for digital image processing.

45. Fluoroscopic examinations with a switched off device for automatic regulation of dose rate are permitted:

45.1. in examining small parts of the body, including extremities;

45.2. when examining a child not older than 5 years or whose weight does not exceed 12 kilograms; or

45.3. in other cases, taking into account the technical possibilities of the radiological installation.

46. Operators shall ensure:

46.1. that fluoroscopic installations are supplied with a device for measuring the patient dose; and

46.2. that an x-ray equipment without the device for automatic regulation of dose rate is not utilised in examinations.

47. Operators shall ensure that the radiological installations to be used in radiotherapeutic procedures and the equipment thereof conform to the following requirements:

47.1. installations and the equipment thereof shall be supplied with devices that provide a choice of medical exposure and the maintenance of operating parameters in conformity with the purpose of use, with continuous and unambiguous indication regarding such technical parameters of the radiological installation as the type of ionising radiation, energy, primary beam modifiers, including filters, distance to the patient, zone dimensions, orientation of the beam, time of medical exposure and the patient dose;

47.2. the source of ionising radiation in radiological installations shall be switched off automatically (the generation of ionising radiation shall be switched off or the source of ionising radiation shall be transferred from working position to a safe keeping position, or the termination or essential reduction of the exposure shall be ensured in some other way) if the supply of electric power is discontinued, and the control mechanism of the beam may be restarted only from the control panel;

47.3. installations of a high dose rate shall be supplied with safety blocking devices or other devices that do not allow to utilise such radiological installations otherwise than indicated on the control panel, and such installations shall be ensured with at least two mutually independent systems for the termination of the operation of the installations;

47.4. in the construction of the safety blocking devices such physical protection systems as codes and keys shall be used, which systems shall guarantee that the operation of radiological installations with a switched off blocking devices when performing repair work or calibration, shall be possible only under the direct control of the operation employees by utilising for the blocking two mutually independent systems that are switched off by two specialists concurrently; and

47.5. control equipment shall without delay give a warning regarding any deflections in the operation of radiological installations and shall terminate the operation of such installations if the deflections exceed the parameters for safe operation.

48. Operators shall ensure that:

48.1. any damage of a component of a radiological installation is determined as soon as possible, but such damage as may create large doses of ionising radiation is determined without delay and, if technically possible, it shall be ensured that the source of ionising radiation is not in working position;

48.2. the probability of a mistake by a human being in the operation of the source of ionising radiation is reduced to the minimum; and

48.3. the technical safety of radiological installations is examined in accordance with the procedures prescribed by regulatory enactments.

49. Radiological installations shall be calibrated when commencing the operation of the source of ionising radiation, at least once a year during the operation, as well as after each operation, maintenance and repair procedures that affect the parameters of ionising radiation, in conformity with the requirements of regulatory enactments and the requirements determined by the manufacturers of the installation.

50. In performing the surveillance of the radiation safety and nuclear safety the centre shall:

50.1. take decisions regarding the possibility to utilise a radiological installation;

50.2. if necessary, require the operator to improve the technical parameters of the radiological installation or to repair defects; and

50.3. require to terminate the utilisation of the radiological installation if the operation of such installation does not conform to the requirements for radiation safety or the technical parameters of the radiological installation do not conform to the prescribed requirements (Annex 8), or the radiological installation does not provide the necessary radiodiagnostic information or radiotherapeutic result.

7. Patient Dosimetry and Evaluation of a Patient Dose

51. In radiotherapy, a medical physicist shall prepare a plan for medical exposure of each patient and shall co-ordinate such plan with the medical physics expert. The plan shall be approved by the practitioner. The following shall be included in the plan for medical exposure:

51.1. the total dose of ionising radiation in the target volume of medical exposure, in the surrounding tissues and in the organs that are sensitive to the ionising radiation (organs of lesser pelvis, mammary glands and thyroid glands);

51.2. the topographical scheme of ionising radiation;

51.3. the source of ionising radiation, the type and energy of ionising radiation, geometrical parameters of the field to be exposed, modifiers of the field, monitor units and other parameters of the field to be exposed; and

51.4. the number of fractions of medical exposure, dose rate of ionising radiation and time period for each fraction of exposure.

52. Practitioners, roentgenology and radiology nurses and medical physicists shall:

52.1. control the performance of the plan for medical exposure of the patient; and

52.2. document the relevant medical radiological procedures in the control card and the register of radiological examinations, certifying it by their signature. A separate control card and register of radiological examinations shall be utilised for each medical radiological procedure.

53. The following information regarding each medical exposure undergone by patients shall be documented in the control card and the register of radiological examinations:

53.1. in radiodiagnostic procedures - the type of examination and the patient dose received during the examination, or:

53.1.1. in fluoroscopy - high-voltage, electric power, the time of medical exposure;

53.1.2. in roentgenography - high voltage, electric power, exposure time and the number of exposures; and

53.1.3. in computed tomography - high voltage, electric power, the number of sections; if a spiral computed tomograph is utilised - high voltage, electric power, the total number of revolutions and the duration of one revolution;

53.2. in radiological procedures of nuclear medicine:

53.2.1. the name of the radioactive pharmaceutical preparation administered;

53.2.2. the total radioactivity of the radioactive pharmaceutical preparation administered; and

53.2.3. date and time of administering a radioactive pharmaceutical preparation;

53.3. in radiotherapeutic procedures:

53.3.1. the organ to be exposed and the target volume;

53.3.2. the total patient dose in the organ and tissues and the dose of ionising radiation per each time of medical exposure;

53.3.3. the dose of ionising radiation in the organs that are sensitive to the ionising radiation (organs of lesser pelvis, mammary glands and thyroid gland);

53.3.4. the dose rate of ionising radiation per each time of medical exposure;

53.3.5. duration of exposure per each time of medical exposure;

53.3.6. the dose rate of ionising radiation outside the body of the patient, to whom a radioactive pharmaceutical preparation has been administered;

53.3.7. topographical scheme of the doses of ionising radiation in the medical exposure; and

53.3.8. the date and time when medical exposure was performed.

54. Operators shall keep the information referred to in Paragraph 53 of these Regulations until the person referred to reaches or should have reached the age of 75 years, but not less than for 30 years after the relevant medical exposure. Upon the request of another medical treatment institution, operators shall provide the requested information regarding the medical exposure undergone by a patient.

55. In establishing and updating the database on population exposure the centre shall:

55.1. compile the information regarding patient doses;

55.2. evaluate the total doses of ionising radiation for separate population groups; and

55.3. analyse the possibilities to reduce the patient dose in conformity with the available technologies and methods.

56. If a medical radiological procedure does not provide the patient with the radiodiagnostic information or the radiotherapeutic result and the patient dose is substantially lower than the diagnostic reference levels, the practitioner shall make the necessary corrections in the performance of medical exposure so that from the subsequent medical exposures sufficient radiodiagnostic information would be obtained or the radiotherapeutic result provided.

57. If the patient dose in the radiodiagnostic procedure exceeds the diagnostic reference level or a person undergoes a non-standard examination, the practitioner shall take measures in order to reduce the medical exposure of the person by utilising this radiodiagnostic procedure henceforward.

8. Protection of Pregnant Women, Breastfeeding Mothers and Children against Ionising Radiation

58. If it is intended for a woman of childbearing age to undergo medical exposure, the prescriber (or practitioner) shall find out whether she is not pregnant or breastfeeding, and shall make a relevant entry in the medical document and referral thereof. The woman shall certify with a signature that the information provided by her is correct.

59. If medical exposure is necessary for a pregnant woman, especially, if abdominal and pelvic regions are subjected to the exposure, the prescriber and practitioner in the justification for medical exposure shall perform additional evaluation of the urgency of the medical exposure and the optimisation of the medical exposure, taking into account also the exposure of the foetus.

60. Any medical exposure of a pregnant woman shall be planned so that she would receive the minimum patient dose. Practitioners shall inform pregnant women regarding the potential risk for the foetus.

61. In the case of a breastfeeding mother, the prescriber and practitioner in the justification for the medical exposure of nuclear medicine shall, in addition, evaluate the urgency of the medical exposure and the optimisation of the medical exposure, taking into account also the expected exposure of the child. Practitioners shall plan the medical exposure so that the mother and the child would receive the minimum patient dose.

62. If in a medical radiological procedure a practitioner utilises radionuclides on a breastfeeding mother, the mother shall terminate the breastfeeding and continue it only when the amount of a radioactive pharmaceutical preparation in the human milk is smaller than the permissible total radioactivity. The permissible limits of the total radioactivity of radioactive pharmaceutical preparations in human milk are specified in Annex 10 of these Regulations.

63. Medical treatment institutions shall organise and support measures where women of childbearing age, pregnant women and breastfeeding mothers shall be warned regarding the potential threats of ionising radiation, and shall promote the awareness of such women by preparing and placing in appropriate places in the relevant medical treatment institutions materials regarding such issues.

64. Practitioners shall ensure that when children undergo medical exposure, only such radiological installations and additional equipment is utilised as is adequate for children, and that the diagnostic reference levels specified in Annexes 2 and 7 of these Regulations are observed.

9. Unplanned Exposure

65. In order to reduce the risk of unplanned exposure, operators shall provide an analysis of potential radiation accidents, ascertain possible damage of radiological installations and mistakes of the employees that could facilitate unplanned exposure of persons.

66. Practitioners shall without delay inform the operators regarding failure to comply with the requirements of these Regulations or of a radiation accident and shall ensure the necessary protection measures during the medical exposure in order to reduce the risk and amount of unplanned exposure of patients.

67. In order to prevent radiation accidents, special attention shall be paid to the technical conditions, the safe operation of radiological installations, to medical radiological procedures, and to the analysis of possible radiation accidents.

68. Operators shall without delay investigate the following radiation accidents and cases of unjustified exposure:

68.1. medical exposure for radiotherapeutic purposes, which exposure has caused harm to the patient if:

68.1.1. inadequate radioactive pharmaceutical preparation has been used;

68.1.2. the dose of ionising radiation for the medical exposure of a person or the fractionation of the dosage differs from that determined by the practitioner;

68.1.3. the medical exposure of a patient may cause unexpected acute effects outside the target volume; or

68.1.4. the total dose of the patient exceeds by 5% the dose determined by the practitioner;

68.2. radio-diagnostic exposure that exceeds by 30% the diagnostic reference level (Annexes 1, 2, 3, 4, 5, 6 and 7);

68.3. unjustified repeated medical exposure of a patient; or

68.4. damage to radiological equipment or other unexpected cases due to which the exposure of patients differs from the planned exposure by more than 20%.

69. If an accident referred to in Paragraph 68 of these Regulations or a case of unjustified exposure has occurred, the operator together with the practitioner and a medical physics expert shall:

69.1. calculate or evaluate the received patient dose and the distribution of such dose in the body of the patient, as well as inform the prescriber and the patient of the dose referred to and the potential consequences of the accident;

69.2. determine and perform a set of correcting measures in order to prevent the recurrence of such accidents;

69.3. within a time period of 24 hours inform the centre of any radiation accidents that may cause serious traumas to a patient or the death of a patient; and

69.4. within a time period of 30 days after a radiation accident which accident may cause serious traumas to a patient or the death of a patient, submit to the centre a written report, in which report the causes of an accident have been specified, information regarding the utilised patient doses, measures taken and any other substantial information provided.

10. Closing Provisions

70. Paragraph 21 of these Regulations comes into force on 1 January 2003.

71. Sub-paragraph 23.2 of these Regulations comes into force on 1 January 2003.

72. Sub-paragraph 43.4 of these Regulations comes into force on 1 January 2004.

73. Sub-paragraph 44.1 of these Regulations comes into force on 1 July 2003.

74. Sub-paragraph 44.2 of these Regulations comes into force on 1 January 2004.

75. Sub-paragraph 46.1 of these Regulations comes into force on 1 January 2006, but with respect to the fluoroscopic installations for children - on 1 January 2003.

76. Sub-paragraph 46.2 of these Regulations comes into force on 1 January 2005.

77. Sub-paragraph 47.4 of these Regulations comes into force on 1 July 2002.

Prime Minister 						A. Bērziņš

Minister for Welfare						A. Požarnovs

Annex 1

Cabinet Regulation No. 97

5 march 2002

Diagnostic Reference Levels in Roentgenography

| No. | Object to be examined | Projections | Entrance surface dose for one patient in one roentgenography (mGy) |
| --- | --- | --- | --- |
| 1. | Lumbar vertebrae | Dorsal (AP) projection | 10 |
|  |  | Lateral (LL) projection | 30 |
|  |  | Lumbosacral transition with an oblique cranially |  |
|  |  | directed beam | 40 |
| 2. | Abdominal cavity | Dorsal (AP) projection | 10 |
| 3. | Pelvic cavity | Dorsal (AP) projection | 10 |
| 4. | Hip joint | Dorsal (AP) projection | 10 |
| 5. | Chest | Front (PA) projection | 0.4 |
|  |  | Lateral (LL) projection | 1.5 |
| 6. | Mammary glands | Craniocaudal (CC) projection (with a grid) | 10 |
|  |  | Mediolateral oblique (MLO) projection (with a grid) | 10 |
|  |  | Lateral (LL) projection (with a grid) | 10 |
| 7. | Urinary system: |  |  |
| 7.1. | before contrasting substance |  |  |
|  | is administered |  | 10 |
| 7.2. | after contrasting substance |  |  |
|  | is administered |  | 10 |
| 8. | Dorsal vertebrae | Dorsal (AP) projection | 7 |
|  |  | Lateral (LL) projection | 20 |
| 9. | Teeth | Intraoral projection | 7 |
|  |  | Dorsal (AP) projection | 5 |
| 10. | Head | Front (PA) projection | 5 |
|  |  | Lateral (LL) projection | 3 |

Note.

1 Dose of ionising radiation into the air (mGy) together with the dispersed radiation at the entrance surface shall conform to the film-screen combination with the relative sensitivity of 200. The value of a film-screen combination of high sensitivity (400-600) shall be reduced two to three times.

Minister for Welfare								A. Požarnovs

Annex 2

Cabinet Regulation No. 97

5 March 2002

Diagnostic Reference Levels in Roentgenography Paediatrics1

| No. | Object to be examined | Projection of entrance surface | dose2 for a patient in one roentgenography (mGy) |
| --- | --- | --- | --- |
| 1. | Chest | Front (PA) projection | 0.10 |
|  |  | Dorsal (AP) projection | 0.10 |
|  |  | Lateral (LL) projection | 0.20 |
|  |  | Dorsal (AP) projection for new-born infants | 0.08 |
| 2. | Head | Front/dorsal (PA/AP) projection | 1.50 |
|  |  | Lateral (LL) projection | 1.00 |
| 3. | Pelvic cavity | Dorsal (AP) projection | 0.90 |
|  |  | Dorsal (AP) projection for infants | 0.20 |
| 4. | Abdominal | Dorsal/front (AP/PA) projection |  |
|  | cavity | with a vertical/horizontal beam | 1.00 |

Notes.

1 The dose of ionising radiation into the air (mGy) intended for a five year old child, in other cases adjustments shall be made by taking into account the age and weight of the patient.

2 The dose of ionising radiation into the air (mGy) together with the dispersed radiation at the entrance surface shall conform to the film-screen combination with a relative sensitivity of 200. The value of a film-screen combination of high sensitivity (400-600) shall be reduced two to three times.

Minister for Welfare								A. Požarnovs

Annex 3

Cabinet Regulation No. 97

5 March 2002

Diagnostic Reference Levels in Computed Tomography

| No. | Type of examination | For one manipulation (many sections) |
| --- | --- | --- |
|  |  | the average dose1 of ionising radiation (mGy) |
| 1. | Head | 50 |
| 2. | Lumbar vertebrae | 35 |
| 3. | Abdominal cavity | 25 |

Note.

1 Determined in the measurements of a phantom of the head produced of a material that is equivalent to water and the length of which along the rotation axis is 15 cm, diameter 16 cm, and of a phantom of abdominal cavity produced of a material equivalent to water, the length of which is 15 cm, diameter 30 cm.

Minister for Welfare								A. Požarnovs

Annex 4

Cabinet Regulation No. 97

5 March 2002

Diagnostic Reference Levels in Mammography

| No. | Average dose of ionising radiation of mammary glands |  |
| --- | --- | --- |
|  | in each craniocaudal projection1 (mGy) |  |
| 1. | With a grid | 3 |
| 2. | Without a grid | 1 |

Note.

1 Determined for a breast that is compressed to a thickness of 4.5 cm and that consists of 50% gland tissues and 50% fat tissues, by utilising an x-ray tube of film and screen system with molybdenum anode and molybdenum filter.

Minister for Welfare								A. Požarnovs

Annex 5

Cabinet Regulation No. 97

5 March 2002

Diagnostic Reference Levels in Fluoroscopy

| No | Type of activities | Dose rate for entrance surface1 |
| --- | --- | --- |
|  |  | (mGy/min) |
| 1. | Fluoroscopy with a normal beam load | 25 |
| 2. | Fluoroscopy with a large beam load2 | 100 |

Notes.

1 Dose of ionising radiation into the air (mGy) together with the dispersed ionising radiation at the entrance surface.

2 Fluoroscopy examinations, the equipment of radiological installations of which provides a possibility to perform fluoroscopy with a large beam load, which is frequently utilised in invasive radiology.

Minister for Welfare								A. Požarnovs

Annex 6

Cabinet Regulation No. 97

5 March 2002

Diagnostic Reference Levels in Nuclear Medicine

| No | Organ to be examined or type of scintigraphy | Radio- nuclide | Chemical form | Maximum total radioactivity of the dosage to be administered per one examination (MBq) |
| --- | --- | --- | --- | --- |
| 1. | Skeleton | 99mTc | Phosphates and other phosphorus | 600 |
|  |  |  | compounds |  |
| 2. | Skeleton (single photon | 99mTc | Phosphates and other phosphorus compounds | 800 |
|  | emission computed |  |  |  |
|  | tomography SPECT) |  |  |  |
| 3. | Red bone | 99mTc | Colloid | 400 |
|  | marrow |  |  |  |
| 4. | Brain | 99mTc | Sodium pertechnetate | 500 |
|  | (static) | 99mTc | Diethylene-triamino-penta | 500 |
|  |  |  | acetate (DTPA), gluconate and |  |
|  |  |  | glucoheptonate |  |
| 5. | Brain | 99mTc | Sodium pertechnetate | 800 |
|  | (single photon emission | 99mTc | Diethylene-triamino-penta-acetate |  |
|  | computed tomography SPECT) |  | (DTPA), gluconate, glucoheptonate | 800 |
|  |  | 99mTc | Exametazime | 500 |
| 6. | Radio-nuclide | 133Xe | Isotonic solution of sodium chloride | 400 |
|  | angiography | 99mTc | Hexa-methyl-propylene amine |  |
|  | of brain |  | oxime (HMPAO) | 500 |
| 7. | Cisternography | 111In | Diethylene-triamino-penta- | 40 |
|  |  |  | acetate (DTPA) |  |
| 8. | Lacrimal glands | 99mTc | Sodium pertechnetate | 4 |
|  | and the drainage thereof | 99mTc | Marked colloid | 4 |
| 9. | Thyroid gland | 99mTc | Sodium pertechnetate | 200 |
|  |  | 123I | Iodide | 20 |
| 10. | Metastases of | 131I | Iodide | 400 |
|  | thyroid gland |  |  |  |
| 11. | Epithelium cells | 201Tl | Thallium chloride | 80 |
| 12. | Lung perfusion | 81mKr | Water solution | 6 000 |
|  |  | 99mTc | Human serum albumin | 100 |
|  |  |  | (macro-aggregates or micro-spheres) |  |
|  |  | 133Xe | Isotonic solution of sodium chloride | 200 |
|  |  | 127Xe | Isotonic solution of sodium chloride | 200 |
| 13. | Lungs (single photon | 99mTc | Human serum albumin | 200 |
|  | emission computed |  | macro-aggregates (MAA) |  |
|  | tomography SPECT) |  |  |  |
| 14. | Lung ventilation | 81mKr | Gas | 6 000 |
|  |  | 99mTc | Diethylene-triamino-penta- | 80 |
|  |  | acetate (DTPA) aerosol |  |  |
|  |  | 133Xe | Gas | 400 |
|  |  | 127Xe | Gas | 200 |
| 15. | Liver | 99mTc | Colloid | 80 |
| 16. | Liver(single photon | 99mTc | Colloid | 200 |
|  | emission computed |  |  |  |
|  | tomography SPECT) |  |  |  |
| 17. | Liver and gall tract | 99mTc | Iminodiacetate derivatives (IDA) | 150 |
| 18. | Spleen | 99mTc | Denatured erythrocytes | 100 |
| 19. | First passage of blood | 99mTc | Sodium pertechnetate | 800 |
|  | through the heart | 99mTc | Diethylene-triamino-penta- | 800 |
|  |  |  | acetate (DTPA), gluconate, |  |
|  |  |  | glucoheptonate |  |
|  |  | 99mTc | Globulin-3 macro-aggregates (MAG3) | 400 |
| 20. | Heart cavities | 99mTc | Human serum albumin | 800 |
|  | and blood vessels | 99mTc | Marked erythrocytes | 800 |
| 21. | Miocard | 99mTc | Phosphonates and phosphate compounds |  |
|  |  | 99mTc | Isonitrils | 300 |
| 22. | Miocard (single photon | 201Tl | Chloride | 100 |
|  | emission computed | 99mTc | Phosphonates and phosphate compounds | 800 |
|  | tomography SPECT) | 99mTc | Isonitrils | 600 |
| 23. | Salivary gland | 99mTc | Sodium pertechnetate | 40 |
|  | and stomach |  |  |  |
| 24. | Meckel's diverticulum | 99mTc | Sodium pertechnetate | 400 |
| 25. | Bleeding of | 99mTc | Colloid | 400 |
|  | stomach and intestinal canal | 99mTc | Erythrocytes | 400 |
| 26. | Emptying of stomach | 99mTc | Non-absorbing compounds | 12 |
|  |  | 111In | Non-absorbing compounds | 12 |
|  |  | 113mIn | Non-absorbing compounds | 12 |
| 27. | Passage of oesophagus and | 99mTc | Non-absorbing compounds | 40 |
|  | regurgitation | 99mTc | Colloid | 40 |
| 28. | Static scintigraphy | 99mTc | Dimercapto-succinic acid (DMSA) | 160 |
|  | of kidneys |  |  |  |
| 29. | Radio-nuclide | 99mTc | Diethylene-triamino-penta- | 350 |
|  | angiography and |  | acetate (DTPA), gluconate and |  |
|  | dynamic |  | glucoheptonate |  |
|  | scintigraphy of kidneys | 99mTc | Globulin-3 macro-aggregates (MAG3) | 100 |
|  |  | 123I | O-iodine hippuran | 20 |
| 30. | Adrenal gland | 75Se | Seleno-cholesterol | 8 |
| 31. | Tumour or abscess | 67Ga | Citrate | 300 |
|  |  | 201Tl | Chloride | 100 |
|  |  | 99mTc | Dimercapto-succinic acid (DMSA) | 400 |
| 32. | Abscess | 99mTc | Exametazime-marked leucocytes | 400 |
|  |  | 111In | Marked leucocytes | 20 |
| 33. | Tumour | 99mTc | Dimercapto-succinic acid (DMSA) | 400 |
| 34. | Neuroectodermal | 123I | Metha-iodine-benzyl-guanidine (MIBG) | 400 |
|  | tumours | 131I | Metha-iodine-benzyl-guanidine (MIBG) | 20 |
| 35. | Lymph nodes | 99mTc | Marked colloid | 80 |
| 36. | Thrombus in blood vessels | 111In | Marked thrombocytes | 20 |

Minister for Welfare								A. Požarnovs

Annex 7

Cabinet Regulation No. 97

5 March 2002

1. Diagnostic Reference Levels in Nuclear Medicine Paediatrics

Table 1

| No | Radio-nuclide | Minimum |
| --- | --- | --- |
|  |  | total radioactivity1 |
|  |  | of dosage |
|  |  | to be administered |
|  |  | to children (MBq) |
| 1. | 67Ga - citrate (tumours) | 10 |
| 2. | 123I - amfetamine (brain) | 18 |
| 3. | 123I - hippuran (kidneys) | 10 |
| 4. | 123I - iodide (thyroid gland) | 3 |
| 5. | 123I - metha-iodine-benzyl-guanidine (MIBG) (tumour) | 35 |
| 6. | 131I - metha-iodine-benzyl-guanidine (MIBG) (tumour) | 35 |
| 7. | 99m Tc - albumin (heart) | 80 |
| 8. | 99m Tc - colloid (liver, spleen and the red bone marrow) | 15 |
| 9. | 99mTc - colloid (stomach regurgitation) | 10 |
| 10. | 99mTc - diethylene-triamino-penta-acetate (DTPA) (kidneys) | 20 |
| 11. | 99mTc - dimercapto-succinic acid (DMSA) (kidneys) | 15 |
| 12. | 99mTc - phosphorus compounds (skeleton) | 40 |
| 13. | 99mTc - denatured erythrocytes (spleen) | 20 |
| 14. | 99mTc - IDA derivatives (gall duct) | 20 |
| 15. | 99mTc - hexamethylpropylene amine oxime (HMPAO) (brain) | 100 |
| 16. | 99mTc - hexamethylpropylene amine oxime (HMPAO) (inflammation) | 40 |
| 17. | 99mTc - human serum albumin macro-aggregates (MAA) |  |
|  | or micro-spheres (lungs) | 10 |
| 18. | 99mTc - globulin-3 micro-aggregates (MAG3) (kidneys) | 15 |
| 19. | 99mTc - pertechnetate (miction cystourethrography) | 20 |
| 20. | 99mTc - pertechnetate (first passage) | 80 |
| 21. | 99mTc - pertechnetate (Meckel's diverticulum |  |
|  | and ectopic mucous membrane of stomach) | 20 |
| 22. | 99mTc - pertechnetate (thyroid gland) | 10 |
| 23. | 99mTc - marked erythrocytes (for the visualising of lumen) | 80 |

Note.

1 1/10 from the total radioactivity intended for adults.

2. Conversion Factors for Children of the Total Radioactivity of Dosage to be Administered to Patients

Table 2

| No | Child's weight | Conversion | No | Child's weight | Conversion |
| --- | --- | --- | --- | --- | --- |
|  | (kg) | factor |  | (kg) | factor |
| 1. | 3 | 0.1 | 15. | 30 | 0.62 |
| 2. | 4 | 0.14 | 16. | 32 | 0.65 |
| 3. | 6 | 0.19 | 17. | 34 | 0.68 |
| 4. | 8 | 0.23 | 18. | 36 | 0.71 |
| 5. | 10 | 0.27 | 19. | 38 | 0.73 |
| 6. | 12 | 0.32 | 20. | 40 | 0.76 |
| 7. | 14 | 0.36 | 21. | 42 | 0.78 |
| 8. | 16 | 0.40 | 22. | 44 | 0.80 |
| 9. | 18 | 0.44 | 23. | 46 | 0.82 |
| 10. | 20 | 0.46 | 24. | 48 | 0.85 |
| 11. | 22 | 0.50 | 25. | 50 | 0.88 |
| 12. | 24 | 0.53 | 26. | 52-54 | 0.90 |
| 13. | 26 | 0.56 | 27. | 56-58 | 0.95 |
| 14. | 28 | 0.58 | 28. | 60-62 | 1.00 |

Minister for Welfare								A. Požarnovs

Annex 8

Cabinet Regulation No. 97

5 March 2002

Minimum Conformity Criteria of Technical Parameters of Radiological Installations

| No | Parameter to be under control | Criteria for requirements/conformity |
| --- | --- | --- |

I. Diagnostic Radiography Equipment1 of General Use1

| 1. | Accuracy of high-voltage: |  |
| --- | --- | --- |
| 1.1. | conformity of the voltage measured | ± 10% |
|  | with the parameters indicated on the scale |  |
| 1.2. | maximum deflection of high-voltage, | ± 10% |
|  | when modifying electric power of the tube |  |
| 1.3. | deflection of high-voltage of the tube for | < ± 5% |
|  | all types of generators in repeated measurements | from average amount |
| 2. | Total filtration of primary beam | 3 2.5 mm Al |
| 3. | Exposure time2 | ± 10% |
| 4. | Dose of ionising radiation: |  |
| 4.1. | dose of ionising radiation with 2.5 mm Al filtration | > 25 µGy/mAs |
|  | at a distance of 1 m from the focus with the voltage of tube of 80 kV |  |
| 4.2. | dose of ionising radiation in repeated exposition | deflection from average |
|  | with the same voltage and filtration | dose ± 20% |
| 4.3. | dose of ionising radiation with the change of | deflection < 15% |
|  | indicated current |  |
| 4.4. | dose of ionising radiation with the change of | deflection < 20% |
|  | indicated current of the tube and exposure time (mAs) |  |
| 5. | Coincidence: |  |
| 5.1. | coincidence of beam and beam of light: |  |
| 5.1.1. | summary deflection in each main axis | £ 3% from the distance |
|  |  | between the focus (radiator) |
|  |  | and the centre of |
|  |  | the visible field (image |
|  |  | detector) |
| 5.1.2. | summary deflection in both main axes | £ 4% from the distance |
|  |  | between the focus (radiator) |
|  |  | and the centre of |
|  |  | the visible field (image |
|  |  | detector) |
| 5.2. | coincidence of the centre of x-ray field and | £ 2% from the distance |
|  | the centre of image detector | between the focus |
|  |  | and image detector |
| 5.3. | coincidence of the centre of the beam of x-rays and | £ 1% from the distance |
|  | the centre of the beam of light | between focus and film |
| 5.4. | coincidence of the centre of beam of light and the film | £ 1% from the distance |
|  | placed in the grid | between the focus and film |
| 5.5. | angle between the axis of the beam of x-rays and | deflection < 1.5o |
|  | the plane of image detector | from 90o |
| 6. | Collimation: |  |

| 6.1. | the beam of x-rays is collimated so that, regardless of |  |
| --- | --- | --- |
|  | the distance between the focus and the image detector |  |
|  | the irradiated field is within the borders of image detector |  |
| 6.2. | in automatic collimation regime the distance | < 2 % from |
|  | between the edges of the beam of light and the edges of image detector in any | the distance3 between |
|  |  | the focus and image |
|  |  | detector |
| 7. | Size of focus field | determined during the whole period of |
|  |  | operation of |
|  |  | the tube4 |
| 8. | Grid: |  |
| 8.1. | defects of image | not permitted |
| 8.2. | obtaining the x-ray image of grid | image is obtained with 50 kV |
| 8.3. | plates of the movable grid shall not be seen in the image |  |
|  | even during the shortest exposure time used in practice |  |
| 9. | Automatic exposure control: |  |
| 9.1. | maximum electric power to restrict | < 600 mAs (except |
|  | over exposure | x-ray radioscopic |
|  |  | examinations and |
|  |  | tomography) |
| 9.2. | restriction of maximum duration of one | 6 seconds |
|  | exposure |  |
| 9.3. | distinction of blackening between two exposures | £ 0.3 for optic |
|  | in the regime of automatic exposure control (AEK), | density units |
|  | one exposure of short and the other of long duration |  |
| 9.4. | maximum difference of optic density of test image | £ 0.3 for optic |
|  | with a fixed thickness of attenuator | density units |
| 9.5. | maximum difference of optic density of test image | £ ± 0.3 for optic |
|  | with a fixed high-voltage if thickness of attenuator | density units |
|  | that conforms to the scope of parameters | from the optic density of |
|  | of a patient is utilised | the average size of |
|  |  | test image |
| 10. | Radiation of the leakage at a distance of 1 m from the focus | < 1 mGy/h |
|  | with a maximum capacity determined by the manufacturer of the tube |  |
|  | that is placed into the casing under examination |  |

II. Development of Films, Characteristics of Image Detector and Conditions for Radiological Image Examination5

| 11. | Intensifying screens and cassettes: |  |
| --- | --- | --- |
| 11.1. | condition and cleanness of the screen and cassettes | on exposed |
|  |  | films no |
|  |  | serious, |
|  |  | important defects |
|  |  | are visible in the image |

| 11.2. | darkening of cassettes - on non-exposed |  |
| --- | --- | --- |
|  | films black edges are not visible if the cassette with |  |
|  | this film has been illuminated at |  |
|  | a negatoscope (each side for 10 minutes) with a luminous |  |
|  | intensity of at least 1000 candlepower/m2 |  |
| 11.3. | contact of film and screen - cassette may not be |  |
|  | a cause for a visible local lack of sharpness or |  |
|  | for a difference in the density of an x-ray picture. The cassette shall be examined by |  |
|  | placing a metal net on the cassette |  |
| 11.4. | relative sensitivity of the film and screen combination of | difference of film |
|  | the same class of sensitivity with | blackening <0.3 |
|  | identical exposure conditions (the same | optic density |
|  | dose of ionising radiation, kilovolts, filtration) | units |
|  | for different cassettes |  |
| 12. | Development of films: |  |
| 12.1. | base and emulsion | < 0.3 optic |
|  |  | density units |
| 12.2. | deflection from the basic value of sensitivity | < 0.2 optic |
|  |  | density units |
| 12.3. | deflection from the basic value of contrast index | < 0.2 optic |
|  |  | density units |
| 13. | Darkened room: |  |
| 13.1. | leak of light is not visible after five minutes long |  |
|  | eye-adaptation in the darkened room with a switched off |  |
|  | inactive and other lighting |  |
| 13.2. | inactive lighting - for a part of film exposed | < 0.1 optic |
|  | in x-ray radiation with the obtained blackening | density units |
|  | for one unit of optic density after four minutes long |  |
|  | lighting at a workplace of the darkened room with |  |
|  | a switched on inactive light and lighting in the neighbouring |  |
|  | rooms, increase of blackening as compared |  |
|  | to the non-lightened part of the same film |  |
| 14. | Conditions for examination of radiological image: |  |
| 14.1. | luminous intensity of negatoscope | 3 1700 candlepower/m2 |
| 14.2. | heterogeneity of light | < 30% |
| 14.3. | lighting of premises | < 50 1x 1 m |
|  |  | distance from |
|  |  | negatoscope |

III. X-ray Radioscopic Examinations6

| 15. | Dose rate: |  |
| --- | --- | --- |
| 15.1. | maximum dose power in the entrance field of | < 0.8 μGy/s |
|  | the intensifier of conventional image with the entrance field diameter of 25 cm |  |
|  | without a grid, exposing an appropriate phantom |  |
|  | (for example, 20 cm thick PMMA phantom) |  |
|  | with automatic control of dose rate |  |
|  | and image brightness control |  |
| 15.2. | utilising a high dose rate, including in invasive | < 1.0 μGy/s |
|  | radiology, the maximum dose rate |  |
| 15.3. | dose rate in cases of other dimensions of | inversely proportional |
|  | entrance field | to the squared diameter |
|  |  | of the field |
| 15.4. | maximum dose rate for a patient or his or her substitute | < 100 mGy/min |
|  | (for example, PMMA phantom) at the entrance |  |
|  | field, including dispersion |  |

| 16. | Resolution of image intensifier and TV chain combination |  |
| --- | --- | --- |
|  | with the following size of the field: |  |
|  | a) 30-35 cm | 0.8 pairs of lines/mm7 |
|  | b) 23-25 cm | 1.0 pair of lines/mm |
|  | c) 15-18 cm | 1.4 pairs of lines/mm |
|  | d) resolution of the image of aimed picture | 2.0 pairs of lines/mm |
| 17. | Contrast threshold in automatic operation regime | £ 4% |
|  | in the image of TV monitor |  |
| 18. | Equipment for termination of operation shall operate automatically, |  |
|  | if the specified summary duration of |  |
|  | x-ray radioscopic examinations (maximum 10 minutes) has |  |
|  | expired. At least 30 seconds before the termination of operation |  |
|  | audio signal is switched on and gives a warning |  |
|  | of the possibility to place the termination equipment in |  |
|  | the starting position, if it is necessary to continue the exposure |  |
| 19. | Cinematography: |  |
| 19.1. | entrance dose rate, when working with an intensifier of | < 0.2 μGy/frame |
|  | 23 cm in diameter |  |
| 19.2. | entrance dose rate for a PMMA phantom of 20 cm | 0.1-0.3 Gy/min |
|  |  | with 25 frames/s |
| 20. | Proportion of the size of entrance field | < 1.15 |
|  | for the field of ionising radiation and image intensifier8 |  |

IV. Conventional Tomography and Computed Tomography

| 21. | Conventional tomography: |  |
| --- | --- | --- |
| 21.1. | difference between the induced and actual | ± 5 mm |
|  | level of the depth of a section |  |
| 21.2. | accuracy of the depth of the section, transferring from one | within the limits of ± 2% |
|  | tomographic plane of the section to the next |  |
| 21.3. | for equipment, the exposure angle of which is larger than | < ± 5° |
|  | 30o, difference between the induced and actual angle9 |  |
| 21.4. | image density for the opening in a lead sheet | homogeneous or |
|  |  | differs in conformity with |
|  |  | the sample that is |
|  |  | intended for the specific |
|  |  | tomography equipment, |
|  |  | image reveals |
|  |  | unexpected depositions, |
|  |  | discrepancies between |
|  |  | exposures or |
|  |  | irregularity of |
|  |  | movement |
| 21.5. | resolution of tomography equipment | 1.6 pairs of lines/mm |
| 22. | Computed tomography: |  |
| 22.1. | HV standard deflection in the centre of examination zone | < 20% of basic |
|  | of the phantom that is equivalent to water or tissues in the are of 5000 mm2 in size | size |
| 22.2. | HV deflection in the phantom that is equivalent to water or tissues | < ± 20 HV |
|  | or in materials of various density at the | or 5% |
|  | same position |  |
| 22.3. | average deflection of numbers in the centre and periphery of the examination zone | < 1.5% in relation to |
|  | of the phantom that is equivalent to water or tissues | the basic size |
|  | in the field of 500 mm2 in size |  |
| 22.4. | measurements of ionising radiation dose index | < ± 20% of |
|  | for one section with any filter and any | the basic size |
|  | depth of the section |  |
| 22.5. | thickness of the layer with a half of the | < 20% of the basic |
|  | ionising radiation dose profile maximum | size |
| 22.6. | for high contrast resolution (spatial resolution) |  |
|  | polystyrene needle image: |  |
| 22.6.1. | difference in density between the maximum and minimum | ± 20% from the basic |
|  |  | value |
| 22.6.2. | difference of geometrical size in relation to | < 20% |
|  | the actual size |  |
| 22.7. | for a low contrast resolution polystyrene needles | visible in image |
|  | with a cross-size of 0.35 cm placed into the phantom |  |
|  | that is equivalent to water |  |

V. Dental Radiology10

| 23. | High-voltage | > 50 kV |
| --- | --- | --- |
| 24. | Filtration of primary beam with a voltage of |  |
| 24.1. | less than 70 kV and 70 kV | equivalent to 1.5 mm Al |
| 24.2. | more than 70 kV | equivalent to 2.5 mm Al |
| 25. | Distance between the focus and skin for radiological |  |
|  | installations with high-voltage: |  |
| 25.1. | of less than 60 kV and of 60 kV | at least 10 cm |
| 25.2. | more than 60 kV | at least 20 cm |
| 26. | Diameter of the size of field at the external end of the tubus | < 60 mm |
| 27. | Timer: |  |
| 27.1. | accuracy | to 20% |
| 27.2. | repeatability | to 10% |
| 28. | Dose of ionising radiation at a distance of 1m |  |
|  | from the focus with the range of voltage 50-70 kV | 30-80 µGy/mAs |

VI. Mammography

| 29. | Generator and control of x-rays: |  |
| --- | --- | --- |
| 29.1. | source of x-rays: |  |
| 29.1.1. | dose rate at a distance that conforms to the distance | at least 7.5 Gy/s |
|  | between the focus and film |  |
| 29.1.2. | distance between the source and image | in conformity with the specifications 3 600 mm |
|  |  | provided by manufacturers |
| 29.2. | coincidence of x-ray field and image detector: |  |
| 29.2.1. | on chest side | x-rays cover |
|  |  | the film 2 5 mm |
|  |  | outside the edges thereof |
| 29.2.2. | on the sides | x-rays cover |
|  |  | the film up to its edges |
| 29.3. | accuracy of high voltage with a range of 25-31 kV | deflection < ± 1 kV |
|  | accuracy of repeatability of high-voltage | deflection < ± 0.5 kV |
| 29.4. | automatic exposure control system: |  |
| 29.4.1. | differences of optic density (including the base and emulsion) | < ± 0,15 optic |
|  | in relation to the standard density | density units |
|  | at a specified point of a developed film11 |  |
| 29.4.2. | differences in density of each level on | from 0.1 to 0.2 |
|  | the scale of the levels of optic density | optic density |
|  |  | units |
| 29.4.3. | deflection of the average size of exposures | < 5% |

| 29.4.4. | differences of the optic density from the optic basic density | ² ± 0.2 optic |
| --- | --- | --- |
|  | obtained for the repeatability of long periods | density units |
| 29.4.5. | dependence of density variation of objects | ± 0.15 optic |
|  | on the thickness of objects | density units |
|  |  | in range relation to |
|  |  | the routine density |
| 29.4.6. | dependence of the optic density variation on | in a range of |
|  | the changes in kV | ± 0.15 optic density |
|  |  | units |
| 29.5. | compression12: |  |
| 29.5.1. | maximum automatic pressure | 130-200 N |
|  |  | (~13-20 kg) |
| 29.5.2. | adjusting of compression plate | minimum |
|  |  | loss of adjustment permitted |
| 29.5.3. | asymmetrical deflection and deflection to the nipple | < 15 mm |
| 29.5.4. | symmetrical deflection | < 5 mm |
| 30. | Grid and image detector: |  |
| 30.1. | exposure factor for the grid system | ² 3 |
| 30.2. | automatic exposure control of the exposure range | within ± 5% |
|  | of all cassettes (mGy or mAs) |  |
| 30.3. | maximum difference in optic density of all cassettes | < 0.2 optic |
|  | in the automatic exposure control (AEK) | density units |
| 31. | Film development: |  |
| 31.1. | base and emulsion | 0.2 optic |
|  |  | density units |
| 31.2. | deflection of sensitivity index from basic value | ± 10% |
| 31.3. | average contrast gradient | > 2.8 |
| 31.4. | daily control of the processor may be performed with sensitometry |  |
|  | at one and the same time approximately one hour |  |
|  | after the starting of the processor. Variability of parameters shall be calculated |  |
|  | after a specified time period, for example, |  |
|  | after a month. Variability of all parameters < 10% |  |
| 31.5. | the darkened room - see Chapter II |  |
| 32. | Negatoscope for the examination of images | luminous intensity 2000- |
|  |  | 6000 candlepower/m2 |
|  | Level of the surrounding light | < 50 lx |
| 33. | Specific features of the system: |  |
| 33.1. | air kerma of entrance surface with: |  |
| 33.1.1. | 40 mm PMMA phantom | ² 10 mGy |
| 33.1.2. | 45 mm PMMA phantom | ² 12 mGy |
| 33.1.3. | 50 mm PMMA phantom | ² 20 mGy |
| 33.2. | for spatial resolution in both directions13 | better than 12 pairs |
|  |  | of lines/mm |
| 33.3. | for contrast measurements of large components with a testing object | < 1.3% contrast |
|  | in a 45 mm thick PMMA phantom, limiting |  |
|  | size for 6 mm components |  |
| 33.4. | exposure time, using a 45 mm PMMA phantom | shorter than two |
|  | to obtain an image | seconds |

VII. Radiotherapy14

| 34. | Gantry rotation | ± 1° |
| --- | --- | --- |
| 35. | Rotation of the head of radiator | ± 2° |
| 36. | Isocentre | ± 2 mm |
| 37. | Indicator of source distance | ± 2 mm |
| 38. | Indicator of beam axis | ± 2 mm |
| 39. | Indication of size numbers of the field | ± 2 mm |
| 40. | Indication of light field | ± 2 mm |
| 41. | Rotation of collimation system | ± 1° |
| 42. | Medical treatment table: |  |
| 42.1. | lateral and longitudinal positioning accuracy | 2 mm |
| 42.2. | vertical positioning accuracy (without a patient) | 2 mm |
| 42.3. | vertical positioning accuracy (with a patient) | 5 mm |
| 43. | Verification system of activities (Gantry angle, | manufacturer's |
|  | size of field, collimator rotation, treatment period or | specifications |
|  | monitor equipment, beam energy) |  |
| 44. | Immobilisation equipment, including blocks, casts | ± 2 mm |
|  | and head, arm and leg clips |  |
| 45. | Equipment for centring of patients | ± 2 mm |
| 46. | Beam parameters and accuracy of the field of light |  |
| 46.1. | indication of the field of light | ± 1 mm for each side |
| 46.2. | calibration of the dose of ionising radiation of the central axis | ± 3% photons |
|  | at the point of interest in a phantom | ± 4% electrons |
| 47. | Stability control: |  |
| 47.1. | 60Co and 137Cs equipment | ± 2% |
| 47.2. | X-ray equipment of a close focus | ± 2% |
| 47.3. | accelerators | ± 2% |
| 48. | Linearity of monitor | ± 1% |
| 49. | Timer of 60Co equipment | ± 0.01 min |
| 50. | Control of electrons/photons | correct type of |
|  |  | ionising radiation |
| 51. | Beam of x-rays: |  |
| 51.1. | homogeneity | ± 3% |
| 51.2. | spatial symmetry | ± 3% |
| 52. | beam symmetry of 60Co and 137Cs equipment | ± 3% |
| 53. | Beam symmetry of an x-ray equipment of a close focus | ± 6% |
| 54. | Homogeneity of electron beams and spatial symmetry | ± 3% |
| 55. | Wedge and compensator transmission factor | ± 2% |
| 56. | Control system of the dose of ionising radiation: |  |
| 56.1. | accuracy | ± 0.5% |
| 56.2. | linearity | ± 1% |
| 56.3. | effect of dose rate | ± 2% |
| 56.4. | stability | ± 2% |
| 56.5. | Gantry angle | ± 3% |
| 57. | Differences in the doses of ionising radiation calculated and | ² 2% |
|  | measured at therapeutically significant points |  |
|  | for computerised distribution of the doses of ionising radiation. |  |
| 58. | Difference of the actual localisation of the isodose curve15 | ² 0.3 cm |
|  | from the calculated localisation. |  |

VIII. Nuclear Medicine

| 59. | Gamma camera |  |
| --- | --- | --- |
|  | (collimator of high resolution - 99mTc): |  |
| 59.1. | homogeneity within the limits of the field used16 | deflection < ± 10% |
| 59.2. | sensitivity (ability to receive gamma radiation | < 20% of basic |
|  | emitted by a radioactive source, impulses/s/MBq) | value |
| 59.3. | deflection of rotation centre (single photon emission | unchanged within the limits |
|  | computed tomography SPECT) | of half-pixel |
| 60. | Multi-head cameras: |  |
| 60.1. | difference in the sensitivity of any head | < 10% |
| 60.2. | offset of pixels of opposing images | < half-pixel |
| 61. | Isotope calibrator: |  |

| 61.1. | linearity | < ± 5% of the range of |
| --- | --- | --- |
|  |  | the activities |
|  |  | used |
| 61.2. | accuracy of repeatability | deflection of ± 5% |
| 61.3. | accuracy of instrument: |  |
| 61.3.1. | for gamma radiators with energy of over 100 keV | deflection of < 5% |
| 61.3.2. | for beta radiators and gamma radiators of | deflection of < 10% |
|  | low energy |  |

Notes.

1 The parameters and criteria referred to in this Chapter apply to the diagnostic radiology equipment for general use. The criteria for the installations of special diagnostic radiology are examined in Chapters III and IV of this Annex. If these chapters do not contain the necessary criteria, the criteria specified in Chapter I shall be utilised.

2 The indicated exposure time > 100 ms.

3 At the same time it is possible to use the field of any size that is smaller than the total size of the image detector field.

4 The deterioration level shall be determined and the further suitability of the tube evaluated.

5 The criteria of this Chapter are intended for obtaining qualitative x-ray images on photo-materials of corresponding quality. In mammography the criteria specified in Chapter VI of this Annex shall be observed.

6 Additional requirements for x-ray radioscopic examinations are formulated in this Chapter. If this Chapter does not contain the necessary criteria, the criteria specified in Chapters I and II of this Annex shall be utilised.

7 Determined by utilising special test objects, for example, grid of Huttner type with 18 resolution levels or Leeds test objects.

8 The situation when collimation borders can be seen on the TV image shall be considered as good practice.

9 For smaller angles this difference is even smaller.

10 The criteria specified in this Chapter are intended for intra-oral radiological installations (or for individual extra-oral radiological installations) but are not intended for the panoramic dental radiological systems. For cephalometric radiological installations the criteria specified in Chapter I may be utilised.

11 Typical standard density shall be from 1.3 to 1.8 optic density units including the base and emulsion.

12 The optimal power of the pressure for pressing the tissues of breasts is not specified, but a special attention shall be paid to the selection of the power of the pressure.

13 Measurements are performed on a testing object that is placed 4 cm above the table (on PMMA surface), on the central line, 6 cm from the edge of the film (on the chest side).

14 The criteria specified in this Chapter are intended for clinical utilisation of radiotherapeutic installations, but are not intended for brachitherapy, intra-operative, dynamic, palliative radiological installations, for radiological installations for the whole body, and for simulators in radiotherapy.

15 In places with very high gradients of the doses of ionising radiation.

16 The test shall be performed with a collimator and without a collimator with the ionising radiation of specified energy (E ± 10%).

Minister for Welfare								A. Požarnovs

Annex 9

Cabinet Regulation No. 97

5 March 2002

Minimum Workload of Medical Physicists and Medical Physics Experts Working with Sources of Ionising Radiation

| No | Source of ionising radiation | Minimum workload of medical physicists1 | Minimum workload of medical physics experts1 |
| --- | --- | --- | --- |
| 1. | Accelerator | 0.88 | 0.37 |
| 2. | Radiological installations |  |  |
|  | (with closed sources of radiation) | 0.34 | 0.14 |
| 3. | X-ray equipment in medicine | 0.07 | 0.03 |
| 4. | Radiological installation utilised for |  |  |
|  | medical exposure in body cavities | 0.42 | 0.18 |
| 5. | Simulator | 0.30 | 0.13 |
| 6. | Radiation planning system: |  |  |
| 6.1. | for external radiotherapy | 0.38 | 0.16 |
| 6.2. | for brachitherapy | 0.08 | 0.04 |
| 7. | For 100 patient years2: |  |  |
| 7.1. | for external radiotherapy | 0.27 | 0.11 |
| 7.2. | for brachitherapy | 0.22 | 0.09 |
| 8. | Activities with radioactive substances |  |  |
|  | (for each TBq) | 0.40 | 0.20 |
| 9. | Portable or mobile |  |  |
|  | radiological installations | 0.30 | 0.15 |
| 10. | Radiotherapy installations, the source of |  |  |
|  | ionising radiation of which in operating condition |  |  |
|  | is located outside the equipment | 0.50 | 0.30 |

Notes.

1 Full workload shall be presumed as 1.

2 100 patient years - apply to patients who undergo the relevant medical radiological procedure or a repeated procedure, or a procedure of a new type (including the procedure of medical exposure if the purpose of exposure is changed) for the first time.

Minister for Welfare								A. Požarnovs

Annex 10

Cabinet Regulation No. 97

5 March 2002

Permissible Total Radioactivity in Human Milk that during a One-year Period is Received by Child

| No | Radio-nuclide | Permissible total radioactivity in a one-year period (MBq) |
| --- | --- | --- |
| 1. | 67Ga | 56 |
| 2. | 99mTc | 5 |
| 3. | 111In | 10.4 |
| 4. | 123I | 0.5 |
| 5. | 131I | 0.01 |
| 6. | 201Tl | 1.2 |

Minister for Welfare								A. Požarnovs